Maxum edition II

Gas Chromatograph Specification

| Configuration |  |
| --- | --- |
| Oven Options | Single isothermal oven or split oven with 2 independent isothermal zones Single or dual independent airless ovens. Dual version has two distinct oven compartments for complete operating independence. |
| Detector Modules | Thermal Conductivity, Flame Ionization, Flame Photometric, Helium Ionization, Electrolytic Conductivity |
| Number of Detector Modules | 1, 2, or 3 in any combination of detector module types for air bath oven 1 or 2 in any combination of detector module types for airless oven |
| Sample/Column Valves | Diaphragm, diaphragm-plunger, and rotating or linear transport |
| Valveless Option | ‘Live’ Switching |
| Columns | Packed, micro-packed, or capillary |
| Gas Supply Regulation | Up to 8 electronic pressure controls and up to 6 mechanical pressure controls |
| Performance |  |
| Minimum Range (general)* | Thermal Conductivity: 0-500 ppm Flame Ionization: 0-1 ppm |
| Repeatability (general)* | ± 0.5% of full scale for full scale ranges from 2-100%; ± 1% of full scale for full scale ranges from 0.05-2%; ± 2% of full scale for full scale ranges from 50-500 ppm; ± 3% of full scale for full scale ranges from 5-50 ppm; ± 5% of full scale for full scale ranges from 0.5-5 ppm; (All values expressed at 2 times standard deviation.) |
| Cycle Time | 15 sec to 3 hr (application dependent) |
| Sensitivity* | ± 0.5% of full scale |
| Linearity* | ± 2% of full scale |
| Oven Temperature Range | 40 to 440 °F (5 to 225 °C) for air bath oven ambient to 535 °F (280 °C) for airless oven |
| Temperature Control | ± 0.05 °F (± 0.02 °C) |
| Ambient Temperature Effect | None with electronic pressure control Varying effect with mechanical pressure control* |
| Vibration Effect | Negligible |
| Mean Time To Repair | 1 hour |
| Mean Time Between Failures* | 3 years excluding consumables |

| Communication Options |  |
| --- | --- |
| Serial Output | RS232, RS485 |
| Ethernet | Standard 10BaseT Ethernet with RJ45 connectors |
| DataNET | Proprietary high speed TCP/IP communication network (redundant pair cable) |
| Data Hiway | Proprietary serial communication network (redundant pair cable) |

| I/O Options |  |
| --- | --- |
| Standard I/O | 2 analog outputs; 4 digital outputs (1 indicates system error, 3 are user configurable); 4 digital inputs; 1 serial output |
| Board Slots for Optional I/O | 2 |
| I/O Boards | AO 8: 8 electrically isolated analog output channels D IO: 4 digital inputs and 4 digital outputs A I/O: 2 digital inputs and 2 digital outputs, 2 analog inputs and 2 analog outputs |
| Digital Inputs | Optocoupler with internal 12-24 Vdc power supply, switch able with floating contacts. Aternative: switchable with external 12-24 Vdc supply (floating relay contacts only), external power supply negative terminal connected to common for specific DI being used |
| Digital Outputs | Floating double-throw contacts, maximum contact load rating 1 A at 30 V |
| Analog Inputs | -20 to +20 mA into 50 ohms or -10 to +10 V Rin = 1 M ohm, mutually isolated to 10 V |
| Analog Outputs | 0/4 to 20 mA into 750 ohms max., common negative pole, galvanically separated from ground, freely connectable to ground |
| Termination | Screw compression terminal strip for braided or solid cable with maximum section of 16 AWG or 1.5 mm2 |
| Gas Sample Requirements |  |
| Sample Flow | 50-200 cc/min (application dependant) |
| Sample Filtration | 5 micron |
| Minimum Sample Pressure | 2 psig (15 KPa), lower pressure optional |
| Maximum Sample Pressure | 75 psig (515 kPa), standard; higher pressure optional |
| Maximum Sample Temperature | 250 °F (120 °C) standard; higher temperature optional |
| Material in Contact with Sample | Stainless steel, Teflon, and polyamide; other material optional |

| Liquid Sample Requirements |  |
| --- | --- |
| Sample Flow | 5-20 cc/min (application dependent) |
| Sample Filtration | 5 micron |
| Minimum Sample Pressure | 5 psig (35 kPa), standard |
| Maximum Sample Pressure | 300 psig (2070 kPa) standard; higher pressure optional |
| Maximum Sample Temperature | 250 °F (120 °C) standard; higher temperature optional |
| Material in Contact with Sample | Stainless steel and Teflon, other material optional |

| Installation |  |
| --- | --- |
| Configuration | Single unit with multiple enclosures |
| Mounting | Wall mount only center to center 44” (1120 mm) Left side clearance: 18" (460 mm) from wall or other equipment Front side clearance: 25 ¾’’ (654 mm) from wall with APU mounted Right side clearance: 18” (460 mm) in all cases |
| Dimensions | Height: 39 ¾” (1010 mm) Width: 26 1/16” (662 mm) Depth: 16 3/16” (451 mm) |
| Weight | 170 lb (77 kg) |
| Enclosure Rating | NEMA 3, IP54 (outdoor with weather protection) |
| EMI/RFI Rating | CE Compliance; certified to 89/336/ECC (EMC directive) CE Compliance; certified to 73/23/EEC (Low Voltage directive) Tested per EN 61010-1 / IEC 1010-1 |
| Hazardous Class | Standard Configurations: Certified by CSA C/US for use in Class I, Division 2, Groups B,C,D Suitable for use in European Zone 2, Group IIB+H2 with local approval Suitable for use in general purpose and non-hazardous areasOptional Configurations: Certified by CSA C/US for use in Class I, Division 1, Groups B,C,D with air or nitrogen purge Certified by CENELEC as EEx pedmib IIB + H2 with air or nitrogen purge and purge control for Zone 1 or Zone 2 |
| Ambient Temperature | 0 to 122 °F (-18 to 50 °C) |
| AC Power | 100-130 VAC or 195-260 VAC (switch selectable), 47-63 Hz., single phase Single oven: 14 amperes max. Dual oven: 2 circuits, 14 amperes max/oven |
| Instrument Air | 50 psig (350kPa) minimum for units using Model 11 or Valco valves 120 psig (825 kPa) minimum for units using Model 50 valves 25 psig (175 kPa) minimum for air bath oven; 3 scfm (85 Lpm)/ oven No instrument air for airless oven |

| Installation |  |
| --- | --- |
| Configuration | Single unit with multiple enclosures |
| Flame Fuel | Hydrogen at 99.999% purity Typical consumption – 70 scf/month/detector module (2000 liters/month/detector module) |
| Flame Air | Zero grade air (< 1ppm THC, O2 content 20-21%). Supplied from instrument air with catalytic purifier (optional). Typical consumption – 900 scf/month (26,000 L/month) |
| Corrosion Protection | Dry air purge to protect electronics Stainless steel oven liner Painted steel exterior (epoxy powder coat) |

| Calibration |  |
| --- | --- |
| Type | Manual or automatic |
| Zero | Automatic baseline correction |
| Span | Standard sample cylinder |
| Mounting |  |

| Notes Unless otherwise specified dimensions are shown as millimeters (Inches) Recommended Clearance 2 Left Side - 460 (18”) Right Side – 460 (18”) Front Side – 654 (25 ¾”) Center to Center – 1120 (44”) Left Exhaust For Single Oven Applications (1”Nipple) Left and Right Exhaust For Split Oven Applications (1”Nipple) |  |
| --- | --- |